STATE OF OHIO DEPARTMENT OF TRANSPORTATION

SUPPLEMENT 1121

CERTIFICATION OF NUCLEAR DENSITY/MOISTURE GAUGES AND OPERATORS January 16, 2026

1121.01	Description.

1121.02	Registration.

1121.03	Operator Certification.

1121.04	Gauge Certification.

1121.01 Description. This Supplement describes the processes and requirements for certification of nuclear density/moisture gauges (“gauge”) and gauge operators for inspection, testing, and/or documentation.

Examination for certification of nuclear density/moisture gauges will be administered by the Office of Materials Management (“OMM”), 1600 West Broad Street, Columbus, Ohio 43223. Examination for certification of gauge operators will be administered by the Office of Construction Administration (“OCA”)

Certified gauges will be listed on the OMM website under Radiation Safety.

Certified Nuclear Gauges | Ohio Department of Transportation

Certified gauge operators will be maintained by the Office of Consultant Services (“OCS”) and will be listed on the OCS website under Individual Prequalification List. https://www.dot.state.oh.us/Divisions/Engineering/Consultant/Pages/individual-preq-list.aspx

1121.02  Registration.

Contact OMM to register and schedule a gauge validation examination. Multiple gauges may be registered to be examined. Gauges should be fully charged, recently calibrated, clean, operational, and contain up to date standard counts on arrival.

The gauge operator certification examination is one requirement for a Soil and Aggregate Inspector in the Consultant Prequalification Requirements and Procedures administered by the OCS. Register for and schedule an examination according to OCA website Consultant Inspection Prequalification Testing.

https://www.transportation.ohio.gov/working/construction/construction-admin/training/inspection-prequal-testing

1121.03  Operator Certification.

At the assigned date, time, and location, candidates must score a minimum of 90 percent correct answers on the written exam. Candidates will have a maximum of 90 minutes to complete the exam. Acceptable materials to take the exam include:

- Provided by the candidate: Pen/Pencil, and a non-programmable calculator.

- Provided by the exam-proctor: the Construction and Material Specifications, the Construction Administration Manual of Procedures, Supplement 1015, Ohio Typical Density/Moisture Curves, Aggregate Correction Chart, and forms CA-EW-5 and CA-EW-6.

1121.04  Gauge Certification.

- Gauge Accuracy Verification.

The Department will verify candidate gauges for use via two methods:

- Field Verification

At the assigned date and time, prepare a test site for the candidate gauge at the location identified by the Department. Perform one-minute direct transmission mode tests at 8-inch and at 12-inch depths for density readings, and a one-minute backscatter mode test for density and the moisture reading. Results from these tests must fall within the acceptable tolerances. The Department will perform testing with the Department’s gauge in the same hole and at the same depths as the candidate’s gauge. The acceptable tolerances, with respect to the Department’s values, are 1) the wet density should compare within ±2.0 pounds per cubic foot and 2) the moisture should compare within ±1.2 percent.

- Shop Verification

At the assigned date and time, bring the candidate gauge to the ODOT Nuclear Shop. Utilizing calibration blocks identified by the Department, perform one-minute direct transmission mode tests at 8-inch and at 12-inch depths for density readings, and a one-minute backscatter mode test for density and the moisture reading. Results from these tests must fall within the acceptable tolerances. The acceptable tolerances, with respect to the Department’s values, are 1) the wet density should compare within ±2.0 pounds per cubic foot and 2) the moisture should compare within ±1.2 percent.

- Certification Period. Gauges will remain certified for 12 months from the verification date unless the calibration is updated, Independent Assurance testing is unfavorable or suspicion of the gauge’s consistent accuracy by the Department is evident. The gauge owner will remedy the gauge’s inaccuracy and prove its accuracy by reverifying the gauge according to 1121.04.A.